## Slide 1: The Christie Proton Research Beamline and Radiobiology Endstation
Mike Merchant
20’ talk 24/06/20
CCAP LhARA Plenary meeting

## Slide 2: The Research Room
The Research Room!
Radiobiology prep room
Varian ProBeam cyclotron
Gantry clinical treatment rooms 1-3
Gantry ‘pit’
4th room and research beamline funded by Christie Charity
Notes: The “research room“ – sounds like a trendy bar?

## Slide 3
Beamline A: Built and commissioned. Final infrastructure works before ready for radiobiology. 20/6/20
Lab access restricted due to COVID-19
Beamline B:
Pre-clinical beamline – in design & evaluation phase
Design aims:
1 mm σ spot
5 cm x 5 cm scanning area
Imaging capability – Xstrahl SAARP
Lots of options - drawing may change!
Water tank interlock
Configurable research space
BPM
Quadrupole triplet
70 – 245 MeV
6 mm σ spot
30 x 40 cm scan area
SAARP

## Slide 4: The Christie Research Beamline
Commissioning tests

## Slide 5: PBT room commissioning
Spot scanning at different energies using Lynx, measuring spot size and scanning speed
Beam line with IBA Lynx to measure beam spots and water phantom to measure Bragg peak (when Lynx removed)

## Slide 6: End-station Requirements
At least 4 separate radiobiology experimental scenarios identified:
Radiobiology
High through-put cell irradiator
  Environmental control (Temperature, CO2, Hypoxia)
  Range of sample types
  Spheroids, Organoids
Single ion tracking cell irradiator (FNTDs, CR-39)
  Low throughput
Bragg peak depth simulator (water tank) – irradiate cells in many different energy layers, energy spectra.
Online microscopy end-station – short term and long term timelapse, observe repair protein kinetics, cell cycle arrest etc.
  Low throughput

## Slide 7: Christie Proton Research Beamline Radiobiology End-station
Varian Engineering Scanning Nozzle
Experimental Trolley
Hypoxia workstation (Don Whitley Scientific)
Beam exit window + lead glass
Cell dish stack (out of beam)
Beam entry window (kapton)
6 axis robot (Thermo Scientific)
245 MeV protons
PTW water phantom
Water tank + Bragg peak detector
Funding for radiobiology end-station from CRUK

## Slide 8: Robotic Automated Hypoxia End-station
Automation for liquid handling: Cell fixing.
FANUC 6-axis robot arm
Proton beam entry window
‘Hotel’ for 36 flasks or well plates
Funding for radiobiology end-station from CRUK

## Slide 9
Proton Hypoxia Radiobiology End-station
Recognised a need to preform high through-put repeatable experiments to reduce experimental error in proton radiobiology.
Wellwash Versa liquid handling station allows us to make best use of imaging facilities in the Manchester Cancer Research Centre.
Version 2: Investigating options for live imaging –
e.g. DNA repair protein kinetics
Funding for radiobiology end-station from CRUK

## Slide 10
Proton Relative Biological Effectiveness
Is there a link between LET and RBE?
Experimental data collated by Paganetti (2014)
Data reproduced from Paganetti, Phys. Med. Biol., 2014
Notes: Equation

## Slide 11: Protons: Mechanistic Modelling of DNA Damage & Repair
Nanodosimetric modelling of DNA damage complexity
Mechanistic modelling of DNA repair
Henthorn et al. Rad. Res. 2017
Henthorn et al. RSC advances 2019
Warmenhoven & Henthorn et al. Nature Scientific Reports 2017
Ingram et al. Nature Scientific Reports 2019
Warmenhoven et al. DNA Repair 2020

## Slide 12: Protons: Biologically Augmented Treatment Planning
Physical Dose
Linear Energy Transfer

## Slide 13: Protons: Biologically Augmented Treatment Planning
DNA Damage model
DNA Repair model
N. Henthorn et al., RSC Advances, 2019
E. Smith et al., Nat. Sci. Rep 2019

## Slide 14
Mechanistic Modelling and Validation of Biological Effect
Measurement
Conceptual Model
Computerised Model
Model Qualification
Model Validation
Model Verification
Iteration
Simulation
Analysis
Development
Clinical Translation
Modelling determines experiments
Multi-assay validation
Equipment designed for experimental need

## Slide 15
Beamline A: Built and commissioned. Final infrastructure works before ready for radiobiology. 20/6/20
Lab access restricted due to COVID-19
Beamline B:
Pre-clinical beamline – in design & evaluation phase
Design aims:
1 mm σ spot
5 cm x 5 cm scanning area
Imaging capability – Xstrahl SAARP
Lots of options - drawing may change!
Water tank interlock
Configurable research space
BPM
Quadrupole triplet
70 – 245 MeV
6 mm σ spot
30 x 40 cm scan area
SAARP

## Slide 16: (Some of) Manchester Research Group in Research Room